## Slide 1: Oscilloscope Basics
What my digital storage oscilloscope can do!

## Slide 2: Table of Contents
Basic Setup of a Digital Oscilloscope
Analogue and Digital Bandwidth
Memory
Signal Sampling
Display
Trigger
Measurement (Cursor, Parameter)
Math
Automatic Signal-Evaluation
Documentation

## Slide 3
Basic Setup

## Slide 4: Basic Setup of a Digital Oscilloscope
ACQUISITION
Input Amplifier Analog/Digital Converter Acquisition Memory Trigger-Logic Time Base
DISPLAY
Display Processor Display Grid
MESSUNG/ANALYSE
Processor Math Co-Processor Processor Memory
DOCUMENTATION
Hard Drive(s) USB-Stick Interfaces
DOCUMENTATION / COMMUNICATION
MEASUREMENT &
ANALYSIS
ACQUISITION
MAIN PROCESSOR
MATH. COPROCESS.
FRONT PANEL
PROCESSOR-MEMORY
PROGRAM &
MATH / ZOOM
SIGNAL DATA
INT. HARD DRIVE
REAL TIME CLOCK
DISPLAY
PROCESSOR DISPLAY
EXT. HARD DRIVE
USB STICK
LAN
RS-232C / GPIB
TIME BASE
TRIGGER LOGIC
EXT.
TRIGGER
CH 1
FAST ACQUISITION MEMORY
ANALOG-
DIGITAL-
CONVERTER
CH 2
FAST ACQUISITION MEMORY
ANALOG-
DIGITAL-
CONVERTER
CH 3
FAST ACQUISITION MEMORY
ANALOG-
DIGITAL-
CONVERTER
CH 4
FAST ACQUISITION MEMORY
ANALOG-
DIGITAL-
CONVERTER

## Slide 5
Bandwidth & Memory

## Slide 6: Analog and Digital Bandwidth
Analog Bandwidth
  Bandwidth of the analog amplifier
Digital Bandwidth
  Depends on the real time sampling rate
  Frequency response of the analog/digital converters
  Memory depth and sampling rate define the time that can be captured
5 MHz rectangle, 250 MS/s
5 MHz rectangle, 2.5 GS/s

## Slide 7: Analog Bandwidth – System
Bandwidth of the analog amplifier
Higher frequencies do not reach the ADC
  Wrong signal shape or no signal display at all
Bandwidth  Probe + Oscilloscope !

## Slide 8: Polling Question #1
What Bandwidth(s) Are You Mainly Using?
  < 100 MHz
  100 – 500 MHz
  500 MHz – 1 GHz
  > 1 GHz

## Slide 9: Memory
Acquisition memory for Signals
  Specification in points ('pts')  e.g. '20 kpts' = 20'000 measured values
Memory per Channel or Interleaved Memory
  Memory combination when few channels are used
   12.5 Mpts per channel (4 channels) / 25 Mpts per channel (2 channels)
Memory Defines the Maximum Sampling Rate for the acquisition time
  1 GS/s sampling at 1 Mpts memory
  1'000'000 measured values / 1'000'000'000 measurements/s = 0.001 s or 1 ms
Memory Segmentation / Sequence Mode
  Fast signals with long signal pauses
  Optimization of memory use

## Slide 10: Memory Segmentation
Efficient and intelligent memory use
Memory division into to 10'000+ segments
Up to 1.25 million acquisitions/s
Capturing fast events with long pauses
Table with time stamp for every segment for clear allocation

## Slide 11
Sampling

## Slide 12: Sampling
Real Time Sampling
  Standard mode in oscilloscopes
  Individual sampling
  Transient signals
  Dependent on the time base and number of channels (interleaved sampling)
  Problem of aliasing
Random Interleave Sampling or Equivalent Time Sampling
  Multisampling
  For periodic signals with a stable trigger time only
  High effective sampling rates are possible
Roll Mode
  Slow signals

## Slide 13: Real Time Sampling
Sine Wave:
  Single frequency
  Sampling rate > 2 x max. frequency
  'Nyquist-Theorem'
Pulse, Rectangular Waves
  Broadband
  Sampling rate > 10 x Max. frequency
Realistically:
  10-fold oversampling

## Slide 14: Aliasing
Reason: Sampling rate too low
Consequence: Frequency too low, wrong waveform, no trigger possible

## Slide 15: Multisampling (Random Interleaved Sampling)
Multisampling of a signal  high effective sampling rate
For periodic signals with a stable trigger time only

## Slide 16: Sampling - Roll Mode
Display of slow processes (e.g. t  0.5 s/div)
The signal passes through the screen from right to left

## Slide 17
Screen

## Slide 18: Screen Display
Acquired Signals
  Direct Display
  Enlarged Display (Zoom, Sections)
  Calculated Signals (Mathematical Combinations)
  FFT Display
Measurements
  Manual Cursor Readout
  Automated Parameter Measurements
  Statistics
Controls / Settings
  Horizontal / Vertical Scale
  Trigger Settings
  Channel Menu

## Slide 19: Screen Display
Display of the Sampling Points or Interpolated Waves
Measured Points
  sampling points, not connected
Interpolated Display
  Linear or Sin x/x connected

## Slide 20: Screen Display
Comparison of Linear Connection and sin (x)/x Interpolation

## Slide 21: Screen Display
Averaging of Measurement Data
Raw Data
  Display of all points
Averaged Representation
  ‘Clean‘ Signal (?)

## Slide 22: Screen Display
Data reduction for display  max/min display
1 Million Data Points
1'000 Screen Columns
  Grouping of the raw data
  Determine of max/min values
  Display the max/min values in the form of a vertical bar in the screen columns
  'Visual noise' by displaying all acquired data points

## Slide 23: Screen Display – Signals
  Menu
  Signals
  Ground
  Time Base / Trigger Channel Info

## Slide 24: Screen Display – Zoom
  Original
  Zoom
  Settings Preview

## Slide 25: Screen Display – Parameter Measurement
  Signal
  Parameter w/ Statistics & Histicons
  Settings

## Slide 26: Screen Display – Cursor Readouts
  Cursors
  Horizontal Values Vertical Values

## Slide 27: Screen Display – Serial Bus Decoder
  Original Signal
  Decoded Contents with Color Overlay
  Table with Values

## Slide 28
Trigger

## Slide 29: Trigger – Overview
Start of Acquisition for Single Events
Synchronizing the Acquisition  Stable Image
Trigger Modes
  AUTO Trigger: continuous triggering on internal time base
  NORMAL Trigger: continuous triggering on defined event (edge, pulse width, logic, smart, …)
  SINGLE Trigger: single triggering on defined event (edge, pulse width, logic, smart, TV, …)
Trigger Sources:
  Analog Channels 1-4
  Digital Channels
  External Trigger Input

## Slide 30: TriggerScan Function
TriggerScan automatically scans the signal for possible triggers
Selection of search criteria
Trigger setting through selection from the compiled list

## Slide 31: Triggering – Standard Triggers
Trigger | Description
 | 'Edge', trigger on positive, negative, or both signal edges
 | Trigger on positive edge
 | Trigger on negative edge
 | Trigger on alternative edges
 | 'Parameter': Trigger on various measurement parameters
 | 'Pulse width': Trigger on positive/negative pulse, defined width/height
 | 'TV Trigger' (PAL, SECAM, NTSC, HDTV 1080i/1080p/720p)

## Slide 32: Triggering – Extended Triggers
Trigger | Description
 | 'Multi-stage Trigger': Event 'A' dependent on 'B' dependent on …
 | 'Interval': Trigger on distance between two same sense edges
 | 'Dropout': Trigger on lack of trigger event within a certain time frame
 | 'Runt': Trigger on reduced positive/negative pulses
 | 'Edge Slope': Trigger when slope outside of definition
 | 'Glitch': Trigger on defined pulse width independent of height
 | 'Window': Trigger when the signal leaves the defined window

## Slide 33: Triggering – Digital Triggers
Trigger | Description
 | 'Digital Pattern': High/Low von CH1, CH2, CH3, CH4 und EXT
 | Logical AND
 | Logical NAND
 | Logical OR
 | Logical NOR
 | Serial Bus Trigger for various protocols

## Slide 34: Triggering – Serial Bus Triggers
Trigger | Description
 | Serial Trigger SPI
 | Serial Trigger I²C
 | Serial Trigger RS-232 / UART
 | Serial Trigger MIL-1553-STD
 | Serial Trigger CAN / CAN FD
 | Serial Trigger LIN
 | Serial Trigger FlexRay
 | Serial Trigger I²S (Audio)
 | Serial Trigger USB

## Slide 35: Polling Question #2
Do You Know And/Or Use These Triggers?
  Slope: Rise/Fall Time Trigger
  Dropout Trigger
  Digital/Logical Trigger Also For Analog Signals
  Serial Data Trigger

## Slide 36
Measurements

## Slide 37: Measurements
Two Measurement Methods:
  Manual cursor readouts
  Automated parameter measurements
Cursor Readouts
  Fast check of all visible points
  Screen resolution only
Parameter Measurements
  Setting of the parameter and/or measurement window
  Full internal resolution
  Statistics

## Slide 38: Measurements – Definition of Horizontal Parameters
Average
(50% Amplitude)
Pulse Width
Pulse Width
Pulse Width
Delay
Period
Period
2 Full Periods
First Point
Last Point
Left Cursor
Right Cursor
Data Area for Surface and Point Calculations, etc.
Trigger Time
Frequency = 1/Period
Frequency = 1/Period

## Slide 39: Measurements – Definition of Vertical Parameters
Maximum
Minimum
Top
Base
Amplitude
Peak/Peak
Pulse Width
Rise Time
Fall Time
Average
(50% Amplitude)
Upper Limit (90% Amplitude)
Lower Limit (10% Amplitude)
Measurement Window, Left Limit
Histogram for Top/Base Identification
Measurement Window, Right Limit

## Slide 40: Measurements - Parameters & Statistics
Automated Online Measurement
Use of the Entire Memory for the Online Parameter Measurement
Display of Statistical Values
'Histicons' to visualize the statistics
Parameters can be used for WaveScan

## Slide 41: Measurements – Parameter Qualification
Qualification of parameters to indicate whether the respective measurement is correct
  'Rise Time': no value due to lack of rising edges
  'Fall Time': possible, but not enough values available on the edge
  'Amplitude': correct measurement

## Slide 42: Polling Question #3
Do You Know/Use Parameter Measurements?
  On A Regular Basis
  Occasionally
  I Prefer Cursor Measurements

## Slide 43
Mathematics

## Slide 44: Mathematics
Conversion of Individual Channels
  e.g. inverting, re-scaling, etc.
Mathematical Combination of Channels
  e.g. 'C1 * C2'
Cascaded Mathematics: f(g(x))
Free definition of customized functions
FFT (Fast Fourier Transformation)

## Slide 45: Mathematics – Customized Functions / WebEdit

## Slide 46: Mathematics – Fast Fourier Transformation FFT

## Slide 47
WaveScan – Automated Signal Monitoring

## Slide 48: WaveScan – Automated Signal Monitoring
Automated Signal Scan
Parameter Monitoring
Search for:
  Runts
  Non-Monotony
  Serial Patterns / Bus
Filter Functions:
  Rarest Events
  Maximum / Minimum
  Within / Outside of Limits, Delta %
Automatic Documentation

## Slide 49: WaveScan – Table of Events
Table with list of found events
Highlighting of the events in the original signal
Selection of a specific event by clicking on the respective table entry
Enlarged display of events

## Slide 50: WaveScan – Search for Digital Patterns
Finding sequences in a digital signal
Display of binary or hexadecimal code
Definition of the sequence to be searched
Table with detected events

## Slide 51
LabNotebook - Documentation

## Slide 52: LabNotebook – Documentation, Storage
Storing of
  the Screen Display
  the Oscilloscope Settings
  the Measured Data of a Channel
  EVERYTHING!  LabNotebook
Saving on
  Internal Memory (Limited)
  Integrated Hard Drive
  External Storage Media (USB
  Drives, etc.)
 Archiving / Further Processing

## Slide 53: LabNotebook – 'Flashback Function'
Reloading of saved Measurements & Settings
Selection of desired LabNotebook archive
Selection of desired entry
Reloading of all data: Analog/Digital Channels, Zoom, Mathematics, Cursor, Parameters, …  Further analyses!
③
①
②

## Slide 54: LabNotebook – Offline Display Using WaveStudio

## Slide 55: LabNotebook – Offline Processing Using WaveStudio

## Slide 56: LabNotebook – Documentation, Reports
Saving setups and all data with a single touch of a button
'Flashback' Function  Retrieving all data and settings of an entry
Report Function  Generating PDF, HTML or RTF reports of the measurement
Emailing the report directly from the oscilloscope

## Slide 57: Questions & Answers
Use the Questions section of the GotoWebinar control panel
Or send your question by email to jens.scheuren@teledynelecroy.com

## Slide 58: Thank You For Your Kind Attention.
Teledyne LeCroy GmbH - Tel. +49 (6221) 8270-0 - contact.gmbh@teledynelecroy.com